# ANNEX II + III :	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

Supply of Digitalization and integration of patients monitoring system in ICU

3955

Columns 1-2 should be completed by the Project partner

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Project partner shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

The requirements set out in the technical specifications represent the minimum technical characteristics which offered goods must satisfy, unless stated otherwise, and tenderers are not allowed to modify technical specification in any way.

For each item for which it is not explicitly stated that it is allowed to offer goods of the equal characteristics, i.e. for each item where it is not stated “or equivalent”, for the purposes of this tender documentation it is assumed that words “or equivalent” are stated, and tenderer is allowed to offer equivalent goods / goods of equivalent characteristics.

| 1. Item no. | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s decision (Y/N) |
| --- | --- | --- | --- | --- |
| 1 | Electric hospital bed with anti-decubitus mattress – 5 pieces Position: Department of Metabolic Disorders - Minimum bed dimensions: length 218, width 99cm - Raising the bed in the range of 39.5 – 76 cm - Bed weight of a maximum of 135 kg - Electrical back adjustment from 0 to 700 - electric adjustment of the foot part from 0 - 350 - electric Trendelenburg adjustment of 150 - A system that extends the mattress platform in the pelvic area and reduces pressure on the tissues in this area when the patient is placed in the cardiac chair position. - additional battery as an alternative power supply - The bed should: Safe working load of minimum 250kg it has 4 horizontal protective plastic point-bumpers on the corners and an additional 2 vertical ones It has an anti-decubitus mattress Dimension 200 x 90 x 14 cm holders on the sides of the bed with hooks from the bag 4 wheels diameter 150mm with 5th central wheel and central locking Plastic fences on all sides of the bed, in the color of your choice with an integrated angle indicator The adjustment is made using the controls located on the outside of the bed, as well as through the manual control panel, which has the options of ACP with trendelenburg + cardiac chair, GO, STOP, TR, aTR, CPR, CPR TR, 4 locks, height adjustment Document holder on the underside of the bed Rails on the sides of the bed Infusion stand, with a minimum of 4 plastic hooks |  |  |  |
| 2 | RESPIRATOR FOR ADULTS AND CHILDREN – 1 piece Position: Department of Metabolic Disorders Technical capabilities: 1 supply of oxygen from the central distribution system of medical gases, with a hose 3 m long 2 air supply using a built-in turbine 3 Ultrasonic flow measurement in the exhalation channel 4 Ability to detect patient disconnection in all NIV modes Ways of ventilation: 7 Volume controlled ventilation (VC) 8 Pressure Controlled Ventilation (PC) 9 Optional possibility of upgrading with pressure-regulated, volume-controlled ventilation (PRVC) 10 Pressure support (PS) with backup ventilation (PC Back-up ventilation) 11 CPAP 12 SIMV combined ventilation modes (SIMV VC+PS, SIMV PC+PS and optional SIMV PRVC+PS) 13 Optional possibility to upgrade with volume supported ventilation ( VS ) mode 14 Optional upgrade with Pressure-controlled ventilation mode at two pressure levels with the permission of spontaneous breaths at both pressure levels (Bi-Vent) 15 Optional possibility to select the mode of automatic switching from pressure, volume or pressure regulated volume controlled ventilation in pressure or volume supported modes in the event of spontaneous breathing (Auto modes) 16 Optional possibility of non-invasive ventilation in pressure controlled (NIV PC) and pressure supported (NIV PS) mode. 17 Optional possibility to upgrade with high flow oxygen therapy (High flow therapy) 18 The possibility of returning to the previous ventilation mode, where all ventilation parameters are saved Ventilation parameters: 19 Adjusting the oxygen concentration from 21 to 100% 20 Inspiratory flow range 0-240 l/min. 21 Inspiratory respiratory volume of 20-2000 ml. 22 Breathing frequency 4-150/min. 23 Minute volume range 0.3-60 l/min 24 Inspiratory pressure minimum 100 cmH2O 25 I:Eratio 1:10 - 4:1 26 The possibility of selecting and adjusting the trigger (trigger) by pressure and flow 27 Flow trigger in the range of 0 - 2 l/min. 28 Pressure trigger in the range of -1 to -20 cmH2O 29 The duration of the apnea alarm is at least 2-45 seconds. 30 Ability to adjust PEEP from 0-50 cmH2O 31 Possibility to choose and set the pause from 0-30 sec. Management monitor: 32 Control monitor with touch-sensitive color LCD screen, minimum dimensions 12 in diagonally, which enables monitoring and setting of all parameters important for mechanical ventilation 33 Automatic adjusting the brightness of the LCD screen 34 A possibility touch screen locks 35 Moving monitor up and down. 36 Minimal display of 3 curves on the ventilator monitor 37 Numerical display of the following values ​​of respiratory parameters - O2 concentration, calculation of inspiratory and expiratory resistance; dynamic and static compliance; display of PEEP; elasticity; Breathing Depth Index (SBI); work of breathing (WOB) of the patient and the ventilator; "P 0.1"; peak, mean and plateau pressure, breathing frequency; end-of-expiry flow; inspiratory and expiratory respiratory volume; inspiratory and expiratory minute volume; spontaneous expiratory minute volume; the relationship between inspiration and expiration; leakage percentage... 38 Optional upgrade with software for visualization of volume and pressure (Peak or Driving) in relation to set target values ​​in invasive modes, with graphic display. 39 Display of trends for a minimum of 72 hours Additional technical features: 40 The possibility of connecting an active humidifier with accessories 42 Movable undercarriage with brakes on all wheels 43 Possibility of use for intrahospital transport 44 Backup power in case of power outage - Internal battery provides at least 120 minutes 45 Communication - USB, LAN 46 Dimensions of the device maximum-489mm x 375mm x 350mm 47 The weight of the device itself is a maximum of 17 kg |  |  |  |
| 3 | Bedside table – 5 pieces Position: Department of Metabolic Disorders - Cabinet for the patient without a feeding accessory. - Possibility to have the cabinet in white color - The cabinet should have a minimum of four double castors with a minimum diameter of 75 mm, for easier manipulation - A minimum of two castors must have the option of locking with a brake - The nightstand must have an integrated side towel holder - The cabinets are accessed from the front, and the drawers must be able to be pulled out in both directions, i.e. from the front and back. - The upper drawer, niche and lower compartment must have sliding doors - Models without integrated feeding board have integrated must have integrated holders for plastic bottles and magazines - The body of the night stand must be made of a painted galvanized cabinet that guarantees high resistance to wear, corrosion and peeling - The top of the night stand is made of laminate resistant to high pressure with extremely high mechanical, chemical and thermal durability. - The edges of the upper side of the bedside table are raised to prevent objects from falling - The dimensions of the fabric must be at least 50x51.5cm - The height of the bedside table must be at least 86.5 cm - The weight of the bedside table must not exceed 35 kg |  |  |  |
| 4 | Hospital bedside table – 5 pieces Position: Department of Metabolic Disorders - Patient cabinet with addition for feeding - The feeding board can be placed on the left or right side of the cabinet without hindering its functionality regardless of where it is placed - The cabinet should have a minimum of four double castors with a minimum diameter of 75 mm, for easier manipulation - A minimum of two castors must have the option of locking with a brake - The cabinets are accessed from the front, and the drawers must be able to be pulled out in both directions, i.e. from the front and back. - The upper drawer, niche and lower compartment must have sliding doors - The body of the bedside table must be made of a painted galvanized cabinet that guarantees high resistance to wear, corrosion and peeling - The top of the nightstand as well as the optional board for feeding the patient are made of high pressure resistant laminate with extremely high mechanical, chemical and thermal durability. - The edges of the upper side of the bedside table are raised to prevent objects from falling - The dimensions of the bed must be at least 50x51.5cm, not including the optional feeding board - The dimensions of the feeding board are at least 60x34.5cm - The height of the bedside table must be at least 86.5 cm - The height of the feeding area must be adjustable in the range of at least 72-108cm - feeding tray must withstand a minimum load of 9 kg - The weight of the bedside table must not exceed 35 kg - Possibility to have cabinet in white color |  |  |  |
| 5 | DEFIBRILLATOR – 1 piece Position: Department of Metabolic Disorders 1. Biphasic, synchronous/asynchronous defibrillator with integrated AED function, 2. Maximum defibrillation energy 200J in steps of: 1-2-3-4-5-6-7-8-9-10-11-12-13-14- 15-30-50-70-90-120- 150-170-200 J., 3. Integrated current/battery supply of the defibrillator, 4. Defibrillator battery capacity 250 shocks with maximum energy or 8 hours of ECG monitoring. 5. The defibrillator has defibrillation pedals with integrated LED indicators to check the patient's resistance, energy selection buttons, charging and shock delivery buttons and a button for printing and registering events, 6. The defibrillation pedals have the option of removing the large defibrillation board for adult resuscitation, which frees the surface of the pediatric pedal for defibrillation of children and automatically activates the pediatric mode, 7. The defibrillator has an integrated color screen, sensitive to touch, with a diagonal of at least 7", protected by tempered glass, 8. The defibrillator has the ability to choose several different screen configurations, 9. The defibrillator has the ability to operate the screen in color and black-and-white "high contrast" mode for operation under direct sunlight., 10. In monitoring mode, the display enables min. 6-channel ECG signal with detection of detached electrodes., 11. The device has an integrated 3-channel thermal printer, with printing of graphic and numerical reports on thermal paper 80 mm wide, 12. Print the report manually at the request of the user and automatically after the electric shock. , 13. The device has a "screenshot" function in all operating modes, 14. The device has an integrated function of automatic external defibrillation with textual and graphic instructions to the operator. 15. The device has an integrated WiFi module for the transfer of recorded data and communication with information systems, 16. The device has the possibility of automatic self-testing of correctness with automatic printing of reports, 17. The device also has the ability to automatically send a report on the performed self-test of correctness via the integrated WiFi module. 18. The device on the back has an integrated hook for hanging the defibrillator on the bed when transporting the patient, 19. Integrated defibrillator memory 24 hours with recording of Defi data, EKG, Impedance curve, registered events, data on performed CPR support, patient data, vital signs and "screenshots" data., 20. Weight of the device without defibrillation pedal max. 5 kg, 21. The defibrillator is supplied with the following standard accessories for operation: - 1x Pedals for defibrillation - 1x Gel for defibrillation - 1x Self-adhesive electrodes for adult defibrillation - 1x 4-wire ECG patient cable - 1x Power cable - 1x Thermal paper - 1x User manual in local language |  |  |  |
| 6 | ECG 12-channel, created according to modern technology – 1 piece Position: Department of Metabolic Disorders 1. 12-channel, high-resolution ECG machine 2. integrated power/battery supply with battery capacity greater than 7 hours of operation 3. ECG sampling frequency min. 30 KHz 4. frequency bandwidth of the device in the minimum range 0-250Hz 5. input impedance not less than 100M Ohm 6. CMRR greater than 90dB 7. the device must have defibrillation protection not less than 5000VDC 8. resolution of the ECG amplifier maximum 1 micro Volt and minimum 24 bits 9. pacemaker signal detection frequency higher than 20KHz 10. integrated color LCD with a minimum resolution of 1024 x 768, a diagonal of no less than 8 inches with the possibility of displaying all 12 drains 11. the device must have a quality check of each drain individually, displayed in different colors depending on the status, the so-called color-coded check 12. the device must to have a system for detecting replaced electrodes when setting 13. the device must have the ability to review the ECG before printing the so-called "Print preview" function 14. integrated memory of the device minimum 350 ECG recordings 15. the device must have the ability to export recordings in PDF format to USB memory 16. the device must have ECG recording with minimum lead configurations: standard, pediatric, right precordial and left posterior. 17. the device must support at least the following formulas for QT correction: Bazett and Fredericia. 18. the device must have the possibility of upgrading "CCAA" or the corresponding software for localization of the place of obstruction in the coronary artery with a risk assessment of the area of ​​the heart at risk and the recommendation of therapy in patients with chest pain 19. the device must have recording of the resting rhythm in duration longer than 9 minutes 20. the device must have an internal 12-channel thermal printer for printing on A4 thermal paper with the possibility of reviewing the ECG before printing, exporting and/or memorization 21. for printing reports, the device must use Z-folded thermal paper with a min width of 210 mm. 22. speed min: 12.5-25-50 mm/sec 23. sensitivity min: 5-10-20 mm/mV 24. the device must have a sealed, rubber or silicone alpha numeric keyboard that is easy to clean and direct function keys 25 . the device must have integrated at least 2 USB interfaces, LAN and WiFi module 26. the device must support two-way communication with hospital information systems. 27. the complete software of the device must be in the local language 28. the weight of the device, maximum 5 kg 29. the device is supplied with standard accessories for work |  |  |  |
| 7 | Infusion stand – 5 pieces Position: Department of Metabolic Disorders The infusion stand should have a minimum of 4 hooks with safety bending The minimum load capacity of the hook must be 2 kg The frame of the infusion stand must be constructed of stainless steel The infusion stand must have a holder with a cup for collecting drops The stand must have at least 5 wheels with a diameter of at least 50 mm At least two wheels must be antistatic A minimum of two wheels must have a locking option The infusion stand must be adjustable in height in the range of at least 135-215 cm The diameter of the base must be at least 63 cm |  |  |  |
| 8 | Infusion pump – 5 pieces Position: Department of Metabolic Disorders Volumetric infusion pump with system pressure measurement: 1 Flow: 0.1-1200ml / h, adjustable 2 The possibility of changing the flow rate without stopping the pump (titration) 3" Accuracy: Typ. +/- 5% 4 "Alarms: visual and audible with text, pre-alarms for status: batteries, system pressure, technical malfunction, incorrect manipulations, dosing. " 5 Bolus flow rate: 10 - 1200 ml / h 6 Continuous infusion pump for intravenous administration 7 Calculation of the drug dose 8 "Adjustable KVO, to maintain the flow of the system for 9 Pump compatibility with photosensitive drug delivery systems 10 "Integrated drug list option: at least 100 drugs in at least 30 categories 11 "Ability to call the last min. 30,000 event stores 12 "Direct power supply at 220 V, with built-in battery with IP33 protection or better " 13 "For better visibility monochrome graphic display, simultaneously displays at least 5 parameters at once - drug, speed, concentration, pressure, remaining time " 14 "With the pump holder installed, the casing condition should allow mounting on horizontal rail and vertical tripod" 15 The software has the possibility of entering the names of medicines composed of at least 20 letters " 16 Have IrDa and RS232 communication port 17 Sound alarm min. adjustable in 3 levels 18 Ability to adjust screen brightness (night lighting) 19 "Battery charging time no longer than 3 hours up to 90% capacity " 20 "Automatic bolus reduction on occlusion release (backoff / anti bolus function) " |  |  |  |
| 9 | SYRINGE PUMP WITH PRESSURE MEASUREMENT IN THE SYSTEM – 5 pieces Position: Department of Metabolic Disorders 1 Flow: 0.1-1200ml / h, adjustable 2 The possibility of changing the flow rate without stopping the pump (titration) 3 Accuracy: Typ. +/- 2% 4 "Alarms: visual and audible with text, pre-alarms for status: batteries, system pressure, syringe, technical failure, incorrect manipulations, dosing. " 5 Bolus flow rate: 10 - 1200 ml / h 6 "Support Syringe Sizes: Automatic detection of syringe volume and size 5ml, 10ml 20ml, 30ml, 50 / 60ml, at least 11 manufacturers" 7 Continuous infusion pump for intravenous administration 8 Calculation of the drug dose 9 Adjustable KVO 10 Pump compatibility with photosensitive drug delivery systems 11 Integrated drug list option: at least 100 drugs in at least 30 categories (Profile) 12 "Ability to recall the last min. 30,000 event stores 13 Direct power supply at 220 V, with built-in battery with IP32 protection 14 Possibility of simultaneous displaying at least 7 parameters at once - drug name, pump speed, drug dose rate, concentration, pressure, remaining time, infusion amount, etc. " 15 "With a built-in pump holder, a derived case condition that allows mounting on a horizontal rail and a vertical tripod " 16 Ability to calculate the administration rate according to The time and volume of the administered drug" 17 The software must enable the entry of drug names consisting of at least 20 letters 18 Sound alarm min. adjustable in 3 levels 19 Ability to adjust screen brightness (night lighting) 20 Battery charging time no longer than 3 hours, up to 90% capacity |  |  |  |
| 10 | Blood warmer and infusion – 2 pieces Position: Department of Metabolic Disorders Unicersal heater of the blood, intravenous fluids and irrigation fluids. The temperature can be set in the minimum range of 35° to 43°C. Temperature can be displayed on a display. |  |  |  |
| 11 | Mattress heater – 3 pieces Position: Department of Metabolic Disorders Mattress heater with temperature ranges from the minimum range of room temperature up to min 39°C. On mobile stand. Supports the use of different shape patient warming mattresses |  |  |  |
| 12 | EEG 60 channels – 1 piece Position: Department of Metabolic Disorders Amplifier 40x EEG channels+ 1x ECG, 2x EOG, 3x EMG, 3x RESP, position sensor, snore sensor, ambient light, Digital channel for patient button - Auxiliary Channel SaO2 (Saturation, heart rate, pulse curve). - Auxiliary Channel etCO2 (EtCO2, Respiratory rate, respiratory curve) 24 bit A\D converter with resolution 3nV/Bit and Input voltage up to 50mV PP sampling rate 256, 512, 1024, 2048 Sps selectable. Input impedance: more than 20GΩ for EEG The common-mode rejection ratio (CMRR) of a differential amplifier > 129 dB, noise less than 1.5 μV pp LEDs display the Impedance at each Input Socket start / stop impedance test via a button on the headbox Photic Stimulator Manual and programmer setup Flash frequency 0-63 Hz Supply voltage 12 VDC Trolley with Isolation Transformer 1000VA 4 anti-static castors with locking brake Isolation Transformer Holder Dimensions: (L) 600 X (W) 560 X (H) 804 mm Keyboard and Mouse Tray Monitor stand Computer CPU i3 or equivalent, 8GB DDR-RAM, 480GB SSD WINDOWS 10 or 11 Professional or equivalent Display 24" FHD high resolution Patient management, Archiving Archiving of patients and EEG / EMG recordings database based on standard Microsoft SQL Pre-defined filters for patients and exams, Standard, ALL examinations, Date (show examination of specific date), Today (all examinations from today) and the last 14 days Easy to customize add filter according to many criteria Software features Automatic Montage change in review Event markers Possibility to add sound to each marker Automatic manual photic stimulation, hyperventilation indication with a timer Recording in split screen showing real time record and recorded EEG Recording in split screen showing part of the record in real time in two different montages Recording & Reviewing with DSA analysis display Sensitivity 1.5, 2, 5, 10, 20, 50, 70, 100, 200, 500, 1000 μV possibility to change this range with unlimited n. of steps Upper cut off frequency filter: 0.1 Hz to 200 Hz customized range/steps with unlimited n. of steps OR any value Time constant: 0.03, 0.1, 0.3, 1, 2s or any value, possibility to use low cut filter Two EEGs of patient records viewing or viewing specific record in 2 different montages in split screen EEG Zoom-Function Available in review mode up to 100X with the following functions: - Possibility to ZOOM with time & Amplitude or separately. -Dynamic mode to adjust the distance between the traces automatically to avoid overlaps between lines The DSA (density spectral array) serves for quick detection of events in the EEG. The DSA representation is a traditional representation mapping the frequencies and spectral power of an EEG signal as a function of time. Spectral analysis is only designed to be used for short artifact-free sections. It is not possible to evaluate the whole EEG with this analysis function. Mapping in four different, preset frequency bands 1.5-3, 3-8, 8-13 and 13-30Hz Spectral parameters of all electrodes are displayed in a table for each selected range. The following parameters can be determined: - Absolute power of the spectral range - Relative power of the spectral range - Median frequency - Edge frequency (95%) - Peak frequency - Mean frequency. |  |  |  |
| 13 | EMG device – 1 piece Position: Department of Metabolic Disorders - TECHNICAL CHARACTERISTICS - electric, visual and acoustic stimulator, 2 trigger-in / 2 trigger-out connectors - Speaker, foot switch - SEP / VEP / AEP - Headphones - Isolation transformer - Trolley - AMPLIFIER - 4 channels with DIN connectors and LED indicator of active channel status - sampling resolution: 24 bits - tester for cables - Upper limit frequency: 20kHz - Notch 50/60Hz filter - Sampling rate: 50kHz/channel - Temperature sensor - Dynamic range EMG ±25Mv - Input impedance > 100MW - CMRR (complete system) > 116dB - Transmission sensitivity 3nV/bit - ELECTRICAL STIMULATOR​ - Operating mode / stimulus strength Constant current, 400 V operating voltage, 0 – 100 mA, adjustable in 0.1 mA steps - Stimulus type:Single stimulus, cyclic stimulus, double stimulus, train - stimulation voltage: 400V - Stimulation frequency 0-100Hz freeselectable - Stimulus series Max. 16 pulses per train with gap and delay - Pulse width 50 μs to 1000 μs - VISUAL STIMULATOR​​ - 19'' monitor - Pattern: checkerboard field, vertical stripes, horizontal stripes - Pattern resolution 4, 8, 16, 32, 64, 128 horizontal fields - Stimulation fields Full field, half field, quarter field - Fixation points Red Dot, Red Cross - ACOUSTIC​​ STIMULATOR - Types of stimulation Click (pressure), clack (suction), alternating, random, noise, sound files with specific frequency and envelopes - Strength of stimulus 0 to 125 dB SPL in 1 dB steps - Load impedance 200Ω - COMPUTER UNIT - Processor at least 8 generations, 8 GB RAM DDR4, SSD, keyboard, mouse, 64 bit professional OS, full hd monitor - SOFTWARE - Anatomical guide - Function for reducing artifacts - EMG: needle EMG, triggers EMG, tremor - Function history stimulation for speeds nervous conductivity - NLG motor, sensory, inching, repetitive stimulation (myasthenia test), F wave, reflexes (H reflex, blink reflex) - 37.The examination program is already preset to record the sensitive action potential after electrical stimulation on a sensitive nerve and to calculate the nerve conduction - velocity (NLG) for two nerves - 38. Function history EMG - 39.Nerve conduction velocity mixed - CTS, PNP, SUS, TTS - SEP - VEP - AEP |  |  |  |
| 14 | Compression Device – 2 pieces Position: Department of Metabolic Disorders 1. System for pressotherapy. 2. Control unit with arm sleeves, abdominal belt and boots. 4. Min. three programs. 5. Rated voltage 220V-240V |  |  |  |
| 15 | Vital functions monitor – 5 pieces Position: Department of Metabolic Disorders - The monitor is intended for continuous monitoring of physiological parameters of patients. - The patient monitor also provides a retrospective review of alarms and parameters from its database. - The monitor must be compact in design and complete with the following functions: - No fan design: No fan on the main unit to eliminate noise and potential risk of infection. - The monitor can work in a barometric environment of 427.5 to 805.5 mmHg (altitude-4550 meters). - 15.6" with minimum screen resolution: 1366 x 768 pixels, with a viewing angle of at least 170 degrees and supporting multi-touch operation. - Automatic brightness: The display must optionally have a light sensor that automatically adjusts the display's brightness to maintain readability in any lighting environment. - Minimum of 12 waveform channels displayed on screen - Waveform display: the bedside display must provide a trace speed of at least 6.25; 12.5; 25 and 50 mm per second. It should allow overlapping of up to 2 IBP waveforms - Clinical measurement data: At the same time, numerical numbers will be displayed including the items below along with the required number of waveforms: - ECG - ST segment - QT/QTc - SpO2/Pulse rate - Breathing rate - Temperature - NIBP - IBP - Side stream/Micro stream/Mainstream CO2 - Up to 26 shortcut keys can be set on the display - Timer: up to 4 timers will be able to be configured on the monitor screen to count down the required time in minutes and seconds with an audio alarm to alert the user. - Online Guide: The monitor will provide a comprehensive online guide to help with quick applications for non-routine clinical measurements. - Operating Modes: The monitor shall provide the following specific operating modes: Standby, CPB, Intubation, Night, Privacy and Demo modes. - Mini trend: The mini trend window can be opened by swiping the screen and the storage duration is from 5 minutes to 8 hours. Any abnormal readings will be highlighted on the trend waveform. Arrhythmia alarm and parameter alarms will be summarized. - Long trend: The trend window can be opened by swiping the screen and the storage duration is from 1 to 24 hours. Any abnormal readings will be highlighted on the trend waveform. Arrhythmia alarm and parameter alarms will be summarized. - Automatic Alarm Limits: Automatic alarm limits for all parameters will be provided so that the patient can be admitted quickly and monitoring started immediately. - False Alarm Suppression: The monitor will be able to analyze multiple parameters to reduce false alarms for heart rate, pulse rate, arrhythmia events. - Infographic alarm indication: a list of technical alarms will be provided, and detailed help messages or images will be available to quickly identify the problem. - Call for help: the system will enable a function on the patient monitor to call nearby medical personnel for help. - Permanent data storage: patient data and measurement records will remain permanently stored in the patient monitor even after power failure and battery discharge. - Full Detection: The monitor will be able to review 48 hours of full detection of all waveforms. - NIBP measurement: up to 3000 groups of NIBP measurements - Early Warning Score (EWS): The monitor will provide an electronic early warning score. - OxyCRG: an apnea event can be identified as an ABD category and record the duration of apnea, nadir HR and SpO2 reading during the event. It can record detailed information during apnea events, such as baby's skin color, stimulation method, and treatment activity. - SpO2 Display: Provides a dashboard to display target period and danger period for SpO2, PR and PI. When the SpO2 reading is higher or lower than the target period, the numeric area will be highlighted. 24 hour SpO2 statistics chart can show percentage of SpO2 target period. - 24 hours ECG summary: supports 24 hours summary of HR, arrhythmias, ST and QT, and the results can be printed by network printer. - NIBP analysis: supports summary for 24 hours of NIBP measurement, and the results can be printed by network printer. - Wireless network: The monitor should provide an integrated Wi-Fi network. - Automatic pacemaker detection: Arrhythmia analysis must be capable of monitoring and analyzing patients with and without a pacemaker. The monitor will be able to detect a pacemaker when the user does not know that the patient has an internal pacemaker. - Arrhythmia classifications in at least 25 arrhythmia classifications, including atrial fibrillation for adult patients. - Option to mount it on an infusion pole or in the wall - The monitor shall be delivered together with an adequate mounting bracket. |  |  |  |
| 16 | Central monitor connected to other monitors – 1 piece Position: Department of Metabolic Disorders Masimo Patient SafetyNet system for remote monitoring of patients' continuous vital signs and timely notification, via a central station, by telephone that enables the display of monitored patients' parameters. Connecting third-party devices must be done via ROOT devices with Iris adapters. It must contain 8 ROOT devices, 5 of which must be connected for remote monitoring via a server connected to a central station and a compatible phone. Sedline modules 4 units, O3 modules 4 units, Radical 7 8 units. Upgradeable for HL7 to transfer data to EMR. |  |  |  |
| NOTE: Service: Provided within 24 hours from the occurrence of the malfunction Training and installation: Included in the price The bidder is obliged to provide proof of the technical characteristics of the offered medical equipment (catalogue, instructions, statement or other proof issued by the manufacturer) |  |  |  |  |