# ANNEX II + III:	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

# Contract title: Supply of LABORATORY EQUIPMENT AGAINST PANDEMIC AND SUPPORTING LABORATORY EQUIPMENT within LAB-OP project, Bar Montenegro

Publication reference: 5912

Columns 1-2 should be completed by the contracting authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the contracting authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words ‘compliant’ or ‘yes’ are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

| 1. Item number | 2. Specifications required | 3. Specifications offered | 4. Notes, remarks, ref to documentation | 5. Evaluation committee’s notes |
| --- | --- | --- | --- | --- |
| LOT 1. EQUIPMENT FOR STERILIZATION AND STERILE MATERIAL DISPOSAL |  |  |  |  |
| 1 | Automatic steam sterilizer (autoclave) 1 Piece 1.1 microprocessor-controlled steam sterilizer with a system for the preparation of water softening with an ion-exchange mass and 3 filters, with a manual (one) door - impassable; 1.2 chamber volume from 60 to 70 liters, 1.3 supply of steam from its own built-in steam generator; 1.4 thermally and sound-insulated chamber doors, manually operated; 1.5 electrical connection 3 phases, 380V, 50 Hz; 1.6 the possibility of sterilizing medical material (textiles, instruments, rubber, plastic) in packed or unpacked form with placement in baskets and boxes, which are then placed via the internal carrier into the chamber of the sterilizer; 1.7 device for sterilization of various medical materials with steam at standard sterilization temperatures of 120-135°C (with sterilization program combinations: a) standard process at 120-135°C; b) express cycle at 134°C; c) the possibility of user defining the process; d) test programs: "Bovie&Dick" test, leak test, etc.); 1.8 water supply from the city water supply; 1.9 sterilizer with injector vacuum pump; 1.10 modern processor control with the possibility of setting individual parameters, computer control of a PC or Laptop, modification of existing sterilization programs, creation of new ones, etc.; 1.11 device management through the control panel; 1.12 the sterilizer must have: a) sound-light signalization for indicating the start/flow/end of the process and errors in the process; b) thermal (alpha-numeric) printer that records the parameter of the sterilization process; and c) safety measures to prevent the door from opening during sterilization, ventilation safety, measuring and control elements, sound and light alarms, etc.; 1.13 chamber and duplicator material: stainless steel; 1.14 material of the supporting structure and device formwork: stainless steel; 1.15 pipe installation and valve material: copper, brass and stainless steel; 1. 16 approximate dimensions of the sterilizer: from 550 to 600 X from 1500 to 1600 X from 850 to 900 mm (width x height x depth); 1.17 Quality Standards ISO 9001, ISO 13485, EN 285:2006. |  |  |  |
| 2 | UPS (uninterruptible power supply device) ONLINE 10kVA, with a set of spare batteries, which are delivered and whose replacement is carried out at the regular annual service 1 Piece 2.1 3 Phase In – 3 Phase Out 2.2 3 Level IGBT Rectifier & Inverter Technology 2,3 Real Digital Signal Processor (DSP) 2.4 High Output Power Factor (PF:1, kVA=kW) 2.5 Increased AC-AC Efficiency (up to 96,5%) 2.6 Wide input voltage range 2.7 Built-in Static & Manual Bypass 2.8 Soft Start Feature 2.9 Parallel connection availability 2.10 Adjustable Battery Qty with optional DC-DC 2.11 Charger/Booster at 10kVA Compact version 2.12 Intelligent battery management system extends the life time of batteries 2.13 Colourful Graphical Multi-Functional TouchScreen LCD Panel 2.14 Event Log Display up to 500 Events 2.15 Advanced communication possibility via RS232 2.16 Generator Port for Generator Friendly Operation 2.17 EPO Port for Emergency Power Off 2.18 Management and monitoring software available for all operating systems 2.19 INPUT 2.19.1 Phase 3Ph+N+PE 2.19.2 Nominal Voltage 380V / 400V / 415V 2.19.3 Voltage Range (%100 Load) (-15%) (+20%) 2.19.4 Voltage Range (%50 Load) (-45%) (+20%) 2.19.5 Nominal Frequency (Hz) 50 2.19.6 Input Current THD* <4% 2.19.7 Input Power Factor 0,99 2.20 OUTPUT 2.20.1 Output Power Factor 1 2.20.2 Phase 3Ph+N+PE 2.20.3 Nominal Voltage 380V / 400V / 415V (adjustable via display) 2.20.4 Static Voltage Regulation @%100 Linear Load <%1 2.20.5 Output Voltage THD* (Online&Battery Mode) < % 2 (Linear Load) 2.20.5 Crest Factor 3:1 2.20.6 Frequency (Hz) 50 Hz 2.20.7 Frequency Range ± %0.01 (Battery Mode) 2.20.8 Overload "Online – Battery Mode: <%125 Load 10 min, <%150 Load 1 min 2.20.9 BypassMode: <%200 continuous’’ 2.20.10 Efficiency up to 96.5% (Online) , 98.5% 2.20.11 Protection Class IP20 2.21 Quality Standards EN 62040-1 (Safety), EN 62040-2 (EMC), EN 62040-3 (Performance), EN 60950. |  |  |  |
| 3 | Set of UPS spare batteries 1 Set for the offered UPS /Item No. 2 of the specification LOT 1/ which are delivered and whose replacement is carried out at the regular annual service. |  |  |  |
| 4 | Medical Ultrasonic Bathtub/Cleaner 1 Piece 4.1 The volume of the device of at least 22 litres; 4.2 Ultrasonic frequency 37kHz (-2/+6) with several working models; 4.3. Setting the working temperature in the range of 40-60 °C; 4.4 Mandatory automatic shutdown. |  |  |  |
| 5 | Closed Type Diesel Electric Generator 1 Piece 5.1 diesel generator set, in the housing for powering the consumer of the facility, operation mode in "STAND BY" power mode from 60 to 70 kVA / from 50 to 60 kW, i.e. from 55 to 65 kVA / from 45 to 55 kW in "PRIME" mode ; 5.2 output voltage 400/230 V, three-phase, 50 Hz, power factor cos φ=0.8, speed 1500 rpm; 5.3 turbo diesel engine, four-stroke, min. 3 cylinders in line, volume min 3,300 L, water cooling; 5.4 fuel consumption at 75% load max 11 l/h; 5.5 tank volume min 145 L; 5.6 max weight with oil and fluids, without fuel from 1000 to 1100 kg; 5.7 digital control panel, must contain a programmable 32-bit microprocessor controller and is designed for automatic as well as manual switching on and off of the generator set; 5.8 "heavy duty" construction, with phosphate chemical pre-protection that ensures corrosion-resistant surfaces; 5.9 instruments: display with adjustable contrast, illumination and auto power off; 5.10 AC measurements: voltage in all 3 phases (L-L & L-N), current per phase and total frequency; 5.11 DC measurements: battery voltage, hours counter, coolant temperature, oil pressure and engine speed; 5.12 engine shutdown protection: low oil pressure, failed start, high engine temperature, too high engine speed, too low engine speed, low or high battery voltage, engine speed loss detection, memory fault tree - possible to read the last 15 to 25 events with the exact time and name of the error; 5.13 controls: RUN-AUTO-STOP buttons with indicators, button for checking the correctness of the lights, emergency shutdown button, menu selection buttons on the display; 5.14 coolant heater, for conditions of safe starting in low temperature conditions; 5.15 automatic control of the condition and charge of the batteries via the charger; 5.16 protective switch; 5.17 shock absorbers to prevent vibrations; 5.18 noise max 63dB at 7m distance at 75% load, supersonic insulated; 5.19 four-pole 100A ATS in its own metal cabinet for automatic switching of consumers from the network to the generator and vice versa; 5.20 delivery according to the turnkey system (with casting of the concrete base, connection to the existing power supply system, with complete electronics necessary for smooth operation); 5.21 Quality Standards: AS1359, CSA C22.2 No100-04, UL142, UL489, UL869, UL2200, NFPA37, NFPA70, NFPA99, NFPA110, IBC, IEC60034-1, ISO3046, ISO8528, NEMA MG1-22, NEMA MG1-33, 2006 /95/EC, 2006/42/EC, 2004/108/EC; 5.22 Exhaust Gas Emission Standard: EU Stage IIIA. |  |  |  |
| 6 | Medical Stainless Steel (Inox) Cabinets 3 Pieces 6.1 Upper half glazed, with shelves; 6.2 Bottom half with sliding doors; 6.3 Material: stainless steel - inox; 6.4 Dimensions: 400*1150*1970mm. |  |  |  |
| 1. Conditions Related to the Trade of Procurement Supplies For Items No. 1 and No. 4 of the specification The subject of the procurement must be registered with the Agency/Institute for Medicines and Medical Devices of Montenegro, which the bidder proves by submitting the Decision on the registration/renewal of the registration of the offered medical device in the registry issued by the Agency/Institute for Medicines and Medical Devices of Montenegro or a certificate of receipt of the request for renewal of the decision on registration in the register of medical devices (with delivery of the decision whose validity has expired), issued by the Agency/Institute for Medicines and Medical Devices of Montenegro. (In the event that the bidder is not the holder of the registration of the offered medical device, he/she is obliged to submit an authorization from the legal person who registered the offered medical device). In case the procurement suplies are not registered in the register of medical devices at the Agency/Institute for Medicines and Medical Devices of Montenegro, the offers of bidders that are registered in the countries of the European Union, the USA, Canada, Switzerland, Norway and the countries of the ex-Yugoslavia will be taken into consideration, for which they must submit the following documentation: a. CE mark depending on the class; b. ISO 13485 quality system certificate; c. Declaration of Conformity; d. Evidence that the medical device is in circulation in the country of manufacture; e. The manufacturer's authorization for the bidder to offer the procurement supplies in the relevant procedure. 2. Other Conditions Related to the Implementation of the Contract 2.1 The bidder is obliged to submit, with the bid, a statement that they will provide authorized service and spare parts in the warranty and post-warranty period of min. 5 years from the delivery of the equipment. 2.2 The bidder is obliged to submit evidence/s in the bid about the technical characteristics of the offered products (catalogue or another document issued by the manufacturer). 2.3 The bidder is obliged to submit a statement that the subject supplies shall be delivered within 15 days at the Contracting Authority premisses. 2.4 The bidder is obliged to submit a statement with their bid that delivery includes: delivery of equipment, installation and testing of equipment, as well as training of personnel and delivery of user instructions in the Montenegrin language. |  |  |  |  |

| 1. Item number | 2. Specifications required | 3. Specifications offered | 4. Notes, remarks, ref to documentation | 5. Evaluation committee’s notes |
| --- | --- | --- | --- | --- |
| LOT 2. LABORATORY EQUIPMENT |  |  |  |  |
| 1 | Laboratory Microscope of Biological Samples 1 Piece 1.1 UIS2 optical system – infinitive optics; 1.2 illumination: minimum 2W LED light source, "Light manager" light intensity control system for each individual lens, Keller illumination; 1.3 coarse and fine focus, fine stroke 0.1 mm/rotation or less, coarse 17.8 mm/rotation or less. The possibility of locking the table in focus; 1.4 adjustment graduation must be 1μm or more; 1.5 manual mount: lens with minimum 5 positions, coded; 1.6 tube trinocular fixed, FN 22, inclination 30° or more, interpupillary distance adjustment in the range of 50-76 mm or more; 1.7 distribution of tube light in at least three positions: 1. 100%-0; 2. 20% - 80%; 3. 0-100%; 1.8 eyepieces: 10x FN 22; 1.9 table: 1.9.1 control on the right, 1.9.2 rectangular coated with ceramic, 1.9.3 moved by wire, 1.9.4 sizes from 150 to 160mm x from 185 to 195mm (movement from 70 to 80 x from 50 to 60 mm), 1.9.5 Double preparation holder, opening from the left side; 1.10 Light condenser: for phase contrast, dark field, bright field minimum NA 1.1; 1.11 Objectives - plan achromatic: 4x (minimum NA 0.1), 10x (minimum NA 0.25), 20x (minimum NA 0.40), 40x phase contrast (minimum NA 0.65), 100xO (minimum NA 1.25), immersion, phase contrast; 1.12 Power supply. |  |  |  |
| 2 | Laboratory Centrifuge with Cooling 1 Piece 2.1 max. RCF 20,913 × g; 2.2 speed: 200–14,000 rpm; 2.3 max centrifuge capacity: min. 4×250 mL/ 2×5 MTP; 2.4 temperature: -9 °C to 40 °C; 2.5 FastTemp function for rapid pre-cooling, with the possibility of several user-defined programs; 2.6 noise level: < 58 dB; 2.7 setting the level of starting/braking; 2.8 swing rotor 1pc: 2.8.1 Max. speed 4,500 × g (5,000 rpm) 2.8.2 Swing bucket rotor with max 4 buckets 2.8.3 max capacity for the rotor: 4x100 ml 2.9 automatic rotor recognition and imbalance detection for maximum operational safety; 2.10 delivered with the device: 2.10.1 adapters for 7 test tubes from 3 – 15 Ml for baskets for rotor 4 pcs, 2.10.2 adapters for 14 test tubes from 1.2 - 5 Ml for baskets for the rotor 4 pcs; 2.11 autoclavable baskets and rotor. |  |  |  |
| 3 | Medical Scale 1 Piece 3.1 capacity 3200g; 3.2 readability 0.01g; 3.3 repeatability 0.001g; 3.4 linearity ±0.02g; 3.5 calibration external; 3.6 stabilization time ~1 s (~0.8 s ¦ 5 g); 3.7 measurement pan dimensions from 160 to 170 × from 160 to 170 mm / from 6,3 to 6,7 × from 6,3 to 6,7 inches; 3.8 tare memory function; 3.9 adjusting the height and tilt of the scale using the scale support; 3.10 Weighing units g, oz, lb, lb-oz, ozt, ct, momme, dvt, gr, com, %, SG, user programmable unit. |  |  |  |
| 4 | Laboratory Refrigerator 1 Piece 4.1 capacity min. 460 litres; 4.2 operating temperatures in the range 0 to +15 °C with display; 4.3 temperature graduation max 0.1°C; 4.4 air convection 4.4 display with microprocessor, internal memory for programs and data storage; 4.5 visual and audible alarm for deviations from the standard operating mode; 4.6 automatic defrosting; 4.7 minimum 11 shelves of stainless steel; 4.8. temperature fluctuation at 4 ° (± ° C) 0.6; 4.9. temperature variation at 4° (±° C) 1. |  |  |  |
| 1. Conditions Related to the Trade of Procurement Supplies For Items No. 1 and No. 2 of the specification The subject of the procurement must be registered with the Agency/Institute for Medicines and Medical Devices of Montenegro, which the bidder proves by submitting the Decision on the registration/renewal of the registration of the offered medical device in the registry issued by the Agency/Institute for Medicines and Medical Devices of Montenegro or a certificate of receipt of the request for renewal of the decision on registration in the register of medical devices (with delivery of the decision whose validity has expired), issued by the Agency/Institute for Medicines and Medical Devices of Montenegro. (In the event that the bidder is not the holder of the registration of the offered medical device, he/she is obliged to submit an authorization from the legal person who registered the offered medical device). In case the procurement suplies are not registered in the register of medical devices at the Agency/Institute for Medicines and Medical Devices of Montenegro, the offers of bidders that are registered in the countries of the European Union, the USA, Canada, Switzerland, Norway and the countries of the fex-Yugoslavia will be taken into consideration, for which they must submit the following documentation: a. CE mark depending on the class; b. ISO 13485 quality system certificate; c. Declaration of Conformity; d. Evidence that the medical device is in circulation in the country of manufacture; e. The manufacturer's authorization for the bidder to offer the procurement supplies in the relevant procedure. For Item No. 3 of the specification The procurement supplies should be registered with the Montenegro Bureau of Metrology. 2. Other Conditions Related to the Implementation of the Contract 2.1 The bidder is obliged to submit, with the bid, a statement that they will provide authorized service and spare parts in the warranty and post-warranty period of min. 5 years from the delivery of the equipment. 2.2 The bidder is obliged to submit evidence/s in the bid about the technical characteristics of the offered products (catalogue or another document issued by the manufacturer). 2.3 The bidder is obliged to submit a statement that the subject supplies shall be delivered within 15 days at the Contracting Authority premisses. 2.4 The bidder is obliged to submit a statement with their bid that delivery includes: delivery of equipment, installation and testing of equipment, as well as training of personnel and delivery of user instructions in the Montenegrin language. |  |  |  |  |

| 1. Item number | 2. Specifications required | 3. Specifications offered | 4. Notes, remarks, ref to documentation | 5. Evaluation committee’s notes |
| --- | --- | --- | --- | --- |
| LOT 3. LABORATORY FURNITURE (Furniture Both for Patients and Laboratory Staff - Patient Chairs, Phlebotomist Chairs, Laboratory Staff Chairs, Computer Chairs, Metal Staff Wardrobe, Table) |  |  |  |  |
| 1 | Blood donor chair 4 Pieces - synchronous adjustable back side and foot with hydraulic system; - blood donor armlet on both sides, it becomes stretcher in emergency; - approximate dimension from 450 to 550 x from 750 to 850 x from 1150 to 1250 mm (W x D x H). |  |  |  |
| 2 | Phlebotomist Work Stool 7 Pieces 2.1 5-star base polished; 2.2 sitting cushion Ø36cm; 2.3 cover in compact upholstered material type eco leather, 2. 4 360° manual ring release; 2.5 safety castors. |  |  |  |
| 3 | Laboratory Staff Work Chair 12 Pieces 3.1 5-star base polished; 3.2 small backrest height adjustable; 3.3 sitting cushion Ø36cm; 3.4 cover in compact upholstered material type eco leather, 3.5 360° manual ring release; 3.6 safety castors. |  |  |  |
| 4 | Work Chair 11 Pieces 4.1 PVC star base with 5 wheels fi50mm; 4.2 height adjustment by gas cylinder; 4.3 PVC armrests, seat and backrest upholstered in furniture fabric. |  |  |  |
| 5 | Three-wing Wardrobe 2 Pieces 5.1 wardrobe according to the HASSP standard; 5.2 made of pickled steel sheet; 5.3 the interior should contain a shelf for storage, a rod for hanging clothes, two hooks per cabinet for hanging clothes; 5.4 the wardrobe must have legs with the possibility of leveling in case of an uneven floor; 5.5 locking with a cylinder lock; 5.6 the doors must have ventilation openings and a tag for the user's name; 5.7 the overall dimensions of the 3-wing wardrobe are approximatelly from 1750 to 1850 x from 850 to 950 x from 450 to 550 mm (H x W x D). |  |  |  |
| 6 | Two-wing Wardrobe 1 Piece 6.1 wardrobe according to the HASSP standard; 6.2 made of pickled steel sheet; 6.3 the interior should contain a shelf for storage, a rod for hanging clothes, two hooks per cabinet for hanging clothes; 6.4 the wardrobe must have legs with the possibility of leveling in case of an uneven floor; 6.5 locking with a cylinder lock; 6.6 the doors must have ventilation openings and a tag for the user's name; 6.7 the overall dimensions of the 2-wing wardrobe are approximatelly from 1750 to 1850 x from 550 to 650 x from 450 to 550 mm (H x W x D). |  |  |  |
| 7 | Chairs in a row, three-seater 15 Pieces 7.1 metal construction made of pipes from 55 to 65 x from 35 to 45 mm; 7.2 plasticized in color; 7.3 seat and backrest plastic or upholstered, 7.4 approximate dimensions: from 1500 to 1600 x 550 to 650 x from 750 to 850 mm (W x D x H). |  |  |  |
| 8 | Ergo table 1 Piece 8.1 plasticized in color; 8.2 metal construction from pipes aproximate from 30 to 40 x from 30 to 40 mm and from 35 to 45 x from 20 to 25 mm; 8.3 worktop raw chipboard cca 20 mm coated and edged with ABS tape; 8.4 approximate dimensions: from 1900 to 2100 x from 850 to 950 x from 700 to 750 (L x W x H). |  |  |  |
| 1. Other Conditions Related to the Implementation of the Contract 1.1 The bidder is obliged to submit a statement that the subject supplies shall be delivered within 15 days at the Contracting Authority premisses. 1.2 The bidder is obliged to submit a statement with their bid that delivery includes: delivery of equipment, installation and testing of equipment, as well as training of personnel and delivery of user instructions in the Montenegrin language. |  |  |  |  |