# ANNEX II + III :	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

Supply of medical equipment for pathology

8470

Columns 1-2 should be completed by the Project partner

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Project partner shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

The requirements set out in the technical specifications represent the minimum technical characteristics which offered goods must satisfy, unless stated otherwise, and tenderers are not allowed to modify technical specification in any way.

For each item for which it is not explicitly stated that it is allowed to offer goods of the equal characteristics, i.e. for each item where it is not stated “or equivalent”, for the purposes of this tender documentation it is assumed that words “or equivalent” are stated, and tenderer is allowed to offer equivalent goods / goods of equivalent characteristics.

| 1. Item no. | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s decision (Y/N) |
| --- | --- | --- | --- | --- |
| 1 | Automated Immunohistochemistry Stainer – 1 piece - System: Fully automatic IHC stainer that performs all IHC processes - Support: Supports IHC, ICC, IF - Capacity: Minimum 30 slides - Chambers: Minimum of 3 separate chambers for 10 slides each - Consumption: 120ul per test - Temperature control and monitoring: Each slide has independent temperature monitoring and control - Reagent container capacity: 7ml and 30ml - Number of containers for reagents: 36 - Volume of reagent box: Minimum 2.5L - Volume of box for toxic waste: Minimum 2.5L - Volume of the external waste box: Minimum 18L - Working temperature: 5-35 degrees - Identification: BarCode identification - Process monitoring: Process monitoring in real time Training and Certification for Immunofluorescence, CISH, and FISH and warranty Mandatory training of a pathologist shall be provided for the use of the system applying Immunofluorescence, Chromogenic In Situ Hybridization (CISH), and Fluorescence In Situ Hybridization (FISH) methods. The training must be conducted in an appropriate accredited reference center and shall include: - theoretical and practical instruction for performing each procedure, - analysis, interpretation, and validation of results in accordance with international diagnostic standards, - issuance of an official certificate upon successful completion of training for each individual method. The total training duration shall be 10 (ten) days per procedure, resulting in a total of 30 (thirty) days for all three methodologies (Immunofluorescence, CISH, and FISH). Upon completion, the participant (pathologist) must be awarded an official certificate confirming competence in performing, analyzing, and interpreting results for Immunofluorescence, CISH, and FISH techniques. The supplier is obliged to provide a minimum two-year (24 months) warranty for the delivered equipment, starting from the date of provisional acceptance and signing of the handover protocol. |  |  |  |
| 2 | Real-Time PCR Detection Systems for In Vitro Diagnostics (IVD) – 1 piece An open system for the use with diagnostic kits and plastic consumables from different manufacturers IVD-CE Fixed block – 96 wells Reaction volume: 1-50 µL Maximum ramp rate: 5°C/sec Average ramp rate: 3.3°C/sec Temperature accuracy: ±0.2oC Temperature uniformity: ±0.3oC Temperature gradient: operational range of 30-100°C, gradient programmable span 1-24°C (on operating range) Sensitivity: 1 copy of target sequence in human genomic DNA Excitation/emission range: 450–684 nm / 515–730 nm Factory calibrations for: Cal Gold 540, Cal Orange 560, Cal Red 610, Cy 5, Cy 5.5, FAM, HEX, Quasar 670, Quasar 705, ROX, Texas 615, VIC, SYBR, Texas Red Other dyes whose excitation and emission ranges match the range of the device can be calibrated via software The device does not require operation with the passive reference color ROX The optical system does not require additional calibration, it is factory calibrated Optical system based on the shuttle system Speed ​​12 s - all channels, FAM/SYBR: 3 s Multiplex: up to 5 targets per well Touch screen USB and Ethernet ports for connectivity, as well as the option to connect to WIFI Software compatible with Microsoft Windows 10 operating system or later versions Ability to connect directly via a USB cable or via a laboratory network via an Ethernet cable 21CFR Part 11 Compliance Portable UV Light Disinfection Lamp Input voltage: 230V/50Hz Life span of UV source: 8000 h IP20 protection Stand on castors: 1760 mm (weight: 2kg) UV Sterilization Cabinet for PCR Preparation, with Spill Tray UV light: 4 x 15 Watt; Nucleic acid inactivation for 5-30 min; timer control White light: 15 Watt; better visibility The door is provided with switches to automatically turn off the UV light when opened Vortex Mixer Speed: 0 - 2850 RPM Working temperature: 4 - 65°C Working mode: touch or continuous Mini Centrifuge complete with 1.5/2.0 ml rotor Speed: 6000 rpm Capacity: 8 x 1.5/2.0ml, 4 x 0.2ml strips Easily interchangeable rotors and adapters, with capacity for 1.5 ml, 0.5 ml i 0.4 ml tubes, as well as 0.2 ml strips and single tubes Max. G: 2000 Spin down time <2 s Ep Research plus G,3-pack, 0.5-10ul,10-100ul,100-1000ul, incl. ep.TIPS boxes 0,5–10 mL, 10–100 mL, 100–1000 mL, variable volume, with stand and 3 boxes of Eppendorf tips 0.5 – 10 µL: systematic error at 0.5 µL: ±0.04 µL, random error at 0.5 µL: ±0.025 µL 10–100 mL: systematic error at 10 µL: ±0.3 µL, random error at 10 µL: ±0.1 µL 100–1000 mL: systematic error at 100 µL: ±3.0 µL, random error at 100 µL: ±0.6 µL Three button work 3-button operation (one operating element for volume selection, one for pipetting, and one for tip ejection) Designed according to the Eppendorf PhysioCare Concept® principles Very low weight and operating forces for advanced ergonomics to limit strain on your hand and arm Temporarily adjustment of pipette in seconds for better accuracy when pipetting challenging liquids like ethanol or when pipetting at high altitudes. Return to factory adjustment without calibration Ability to autoclave the entire pipette or only the lower part according to needs and applications to ensure a proper decontamination High UV and chemical resistance ISO 8655 Eppendorf ThermoMixer® F2.0 or equivalent Power source: 220 – 240 V, 50 – 60 Hz Max. speed: 1,500 rpm Temperature range: Min: 4 °C above room temp., max: 100 °C Mixing frequency:300 – 1,500 rpm Mixing and vortexing radius: 1.5 mm Temperature settings: 1 °C / 100 °C Temperature accuracy: Max. ±0.5 °C at 20 – 45 °C Heating rate: Max. 13 °C/min Anti-spill technology prevents lid wetting and cross-contamination Predefined temperature keys: 37 °C, 42 °C, 56 °C, i 95 °C Centrifuge MiniSpin® plus, with rotor or equivalent Max. RCF: 14100 × g Max. RCF with rotor: 14100 × g Automatic rpm/rcf conversion Speed: 800 – 14,500 rpm Max. capacity: 12 × 1.5/2.0 mL Separate short-spin key Acceleration time: 13 s Deceleration time: 12 s Timer: 15 s - 99 min Noise level: <52 dB(A) Height (with open lid): 25 cm non-refrigerated Automatic lid opening at the end of the run to prevent sample warming and to allow easy access to samples Metal rotor bowl Training and Warranty Mandatory training of a pathologist shall be provided for a duration of 15 (fifteen) days, to be conducted after delivery and installation of the system. The training must be carried out in an appropriate accredited reference center and shall include: - operation and application of Real-Time PCR Detection Systems for In Vitro Diagnostics (IVD), - analysis, interpretation, and validation of obtained results in accordance with international standards. Upon completion of the training, the participant (pathologist) must be awarded an official certificate confirming competence in operating, analyzing, and interpreting results generated by the Real-Time PCR IVD system. The supplier is obliged to provide a minimum two-year (24 months) warranty for the delivered equipment, starting from the date of provisional acceptance and signing of the handover protocol. |  |  |  |
| 3 | Digital slide scanner – 1 piece Numerical aperture: minimum 0.80 Resolution, microscope: minimum 0.25 μm/pix Depth of focus: minimum 1 μm Focusing: Electronic - coarse and fine. Image sensor: minimum 12 Mpix Slide capacity: minimum 1 slide Slide format: 75 x 25 mm Scan speed: 15mm x 15mm area in 3 minutes or less, with resolution of 0.25 μm/pix or better Connectivity: 1 GigE, 802.11 ac WLAN Image formats: TIFF, SVS Illumination: Koehler, LED Operated via a web-interface from desktop, laptop or tablet, no need for software instalation. Local and remote control of all functions—scanning, live image, and scan browsing. Scanning slides 75x25mm with brightfield ilumination and continous autofocus. Unit is delivered with carry-on luggage. Wireless connectivity for scanner control and image sharing. Live images can be accesed and navigated remotely. All data and images are stored onboard the scanner. Warranty The supplier is obliged to provide a minimum two-year (24 months) warranty for the delivered equipment, starting from the date of provisional acceptance and signing of the handover protocol. |  |  |  |
| 4 | Fluorescence upgrade of current microscope system model BX43 Olympus – 1 piece Plan apochromat objective: Magnification 2x, with minimum NA 0.08, working distance 6.2 mm Plan fluorite (plan semi apochromat) objectives: Magnification 4x, with minimum NA 0.13, working distance 17mm Magnification 10x, with minimum NA 0.30, working distance 10mm Magnification 20x, with minimum NA 0.50, working distance 2.1mm Magnification 40x, with minimum NA 0.75, working distance 0.51 mm Magnification 100x, oil immersion, with minimum NA 1.30, working distance 0.2mm 5 x Fluorescence filter cubes: Exciter filter: 371 - 399 nm, barrier filter: 416 - 480 nm. Exciter filter: 426 - 446 nm, barrier filter: 460 - 500 nm. Exciter filter: 472.5 - 497.5 nm, barrier filter: 510 - 540 nm Exciter filter: 532 - 554 nm, barrier filter: 576 - 596 nm. Exciter filter: 567.5 - 592.5 nm, barrier filter: 610 - 640 nm. Coded Fluorescence illuminator with minimum 8 positions for filter cubes LED fluorescence excitation light source, minimum three LED channels with excitation spectrum from 365-635nm (DAPI – Cy5) or wider. Global light intensity control from 0-100%, (1% increments) via control pod. Warranty The supplier is obliged to provide a minimum two-year (24 months) warranty for the delivered equipment, starting from the date of provisional acceptance and signing of the handover protocol. |  |  |  |
| 5 | Tissue Cassette Printer – 1 piece - Contactless UV laser printer - Print: Symbols, numbers, 1D and 2D barcodes - Print speed: up to 3 seconds - Print resolution: Minimum 2500 DPI - Screen: Touch sensitive - Print resistance: Print resistant to chemicals, heat and scratches - Cassettes: Supports cassettes with covers, without covers, with angles of 35-45 degrees - Filter: Built-in air purification system, carbon filter - Magazine capacity: up to 80 cassettes - Remote printing and print monitoring, two-way LIS communication - Support: Supports LIS, HIS, PIC, PACS - Operating system: Windows 7, 10, 11 and Linux Warranty The supplier is obliged to provide a minimum two-year (24 months) warranty for the delivered equipment, starting from the date of provisional acceptance and signing of the handover protocol. |  |  |  |
| 6 | Glass Slide Printer – 1 piece - Contactless UV laser printer - Display: Touch screen - Total capacity: 200 slides - 2 baskets of 100 slides each - Double input and double output of slides - Print speed: Up to 5 seconds per slide - Print resolution: Minimum 2500 DPI - Technology: Technology of pulling the slides from the top - Printing: alphanumeric, various languages, symbols, 1D and 2D barcodes, graphically designed templates - Filter: Built-in triple carbon filter for air purification - Remote printing and print monitoring, two-way LIS communication - Support: Supports LIS, HIS, PIS, PACS - Ports: USB port and ethernet - Operating system: Windows 7, 10 and Linux Warranty The supplier is obliged to provide a minimum two-year (24 months) warranty for the delivered equipment, starting from the date of provisional acceptance and signing of the handover protocol. |  |  |  |